## READ ME
READ ME File for COMET (Counting Method Evaluation Tool)
https://cell-counting.shinyapps.io/COMET/
COMET (COunting Method Evaluation Tool) is a Shiny App that accepts data in a specific template based on a specified experimental design and conducts statistical analysis, and outputs quality indicators for the cell counting methods as well as additional meta-data important for reporting of results.
COMET is designed to accept data from experimental designs that follow the principles outlined in the ISO standard, "ISO 20391-2:2019 Biotechnology — Cell counting — Part 2: Experimental design and statistical analysis to quantify counting method performance". The outputs of this app are intended to follow the reporting requirements and recommendations of the standard.
This excel file contains:
| 1) | Full Data Template for COMET |  | Comprehensive data template with columns for meta-data, that can be saved as a *.CSV file for upload into COMET
| 2) | Full Template Descriptions |  | Description of the columns in the full data template
| 3) | Simple Data Template for COMET |  | A simple form data template with minimal data inputs for running COMET, that can be saved as a *.CSV file for upload into COMET
| 4) | COMET Input Parameters |  | Descriptions of the input parameters for COMET
| 5) | COMET Output Tabs |  | Descriptions of the output tabs and their contents
| 6) | Download Results File Descr. |  | Descriptions of the contents of the file generated by the "Download All Results" button
| 7) | Data Template Examples |  | Example studies that can be saved as *.CSV files and run using COMET
|  |  |  | Full Template Example 1: | Data corresponds with ISO 20391-2:2019 Annex D (Use Case 1)
|  |  |  | Full Template Example 2: | Data corresponds with ISO 20391-2:2019 Annex E (Use Case 2)
|  |  |  | Full Template Example 3: | Data corresponds with publication, Sarkar et. al., Cell & Gene Therapy Insights, 2019
|  |  |  | Simple Template Example 1: | Independent data set
More information on the experimental design and statistical analysis approach for quantifying cell counting method performance can be found at:
https://www.nist.gov/programs-projects/evaluating-quality-cell-counting-methods-experimental-design-and-statistical
Additional Resources:
| 1) | ISO 20391-1:2018 Biotechnology — Cell counting — Part 1: General guidance on cell counting methods | [6 empty] | https://www.iso.org/standard/68879.html
| 2) | ISO 20391-2:2019 Biotechnology — Cell counting — Part 2: Experimental design and statistical analysis to quantify counting method performance | [6 empty] | https://www.iso.org/standard/67892.html
| 3) | Standards Landscape in Cell Counting: Implications for Cell & Gene Therapy. Sarkar, Pierce, Lin-Gibson, Lund. (Cell & Gene Therapy Insights, 2019) | [6 empty] | https://insights.bio/immuno-oncology-insights/journal/article/236/standards-landscape-in-cell-counting-implications-for-cell-gene-therapy
| 4) | Evaluating the quality of a cell counting measurement process via a dilution series experimental design. Sarkar, Lund, Vyzasatya, Vanguri, Elliott, Plant, Lin-Gibson.(Cytotherapy, 2017) | [6 empty] | https://pubmed.ncbi.nlm.nih.gov/29037942/
|  |  | Example Experimental Design |  |  | Quality Indicator Reporting |  |  |  | Meta-Data Reporting
[7 empty] | Ydfi | mean cell count for a set of ni representative test samples, with target dilution fractions dfi
[7 empty] | %CVdfi | mean percent CV for a set of ni representative test samples, with target dilution fractions dfi
[7 empty] | R2 | coefficient of determination
[7 empty] | PI | proportionality index, calculation is method specific. Several calculations for PI can be found in the Cell Counting Part 2 documentary standard

## Full Data Template for COMET
counting_method | target_dilution_fraction | measured_dilution_fraction | random_sample_number | rep_obsv | cell_conc | raw_count | percent_viable_cells | time_elapsed | cell_type | concentration_type | stock_solution | starting_soln_conc | stock_extraction | replicate_sample | analyst
|  | NA |  |  |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA

## Full Template Descriptions
NOTE: Column names must be used exactly. Column order can be rearranged.
Column Name | Date Required | Column Description | Input Type | Default Value | Notes
counting_method | Yes | Label identifying a particular counting method | Character String (e.g. Hemocytometer 1) | "Method 1" | The counting method referes to the measurement process used for counting the cells.
target_dilution_fraction | YES | targeted/desired dilution fraction value | Fractional Number (e.g. 0.3) |  | The target dilution fraction (DF) is the ratio between the expected volume of stock solution to the total expected volume of the sample (expected volume of stock solution + expected volume of diluent)
measured_dilution_fraction | NO | measured value of dilution fraction based on a secondary traceable method | Fractional Number (e.g.0.285) |  | The measured dilution fraction (DF) is the ratio between the measured volume of stock solution to the total measured volume of the sample (measured volume of stock solution + measured volume of diluent). Measured volume may be obtained via a secondary traceable method, for example via weight verified pipetting. If a measured DF is not obtained, "NA" may be entered
random_sample_number | NO | Numerical value identifying the sample | Integer (e.g. 1) |  | Generally, the sample number will be a random number assigned to the sample to identify the sample while keeping the actual sample composition blinded to the analyst. Typically samples are analyzed in the order of the assigned random numbers.
rep_obsv | YES | Numerical value identifying the particular observation of the sample replicate | Integer (e.g.1) |  | The experimental design specifies that replicate observations should be taken for each replicate sample at each target DF. This numerical value identifies which observation of the replicate sample the data corresponds to.
cell_conc | YES | Measured cell concentration for each observation (cells/mL) | Rational Number (e.g. 145600) |  | Cell concentration data from the counting method corresponding to the concentration type. Cell concntration should be expressed as cells/mL
raw_count | NO | Measured raw cell count for each observation (# of cells) | Integer (e.g.435) |  | If available, raw cell count data can be reported here. If not available "NA" may be entered
percent_viable_cells | NO | Measured cell viability (%) | Intiger (e.g. 87) |  | If available, % cell viability data can be reported here. If not available "NA" may be entered. The Shiny App will output analysis of cell viability if this information is provided
time_elapsed | NO | Time (min) elapsed between the time samples were prepared and the time the cell count observation was obtained | Rational Number (e.g. 24.5) |  | Time elapsed is used to evaluate any temporal bias in the data set. To calculate time elapsed the time at which the final sample was prepared should be recorded, and the time at which each observation was made should be recorded. Time elapsed for each observation can then be indicated as the number of minutes elapsed between these two events. If time elapsed data was not collected "NA" may be entered.
cell_type | NO | Label identifying a particular cell type or preparation | Character String (e.g. Fresh NIH 3T3) | "Cell Type A" | The cell type refers to the type of cell being enumerated or the specific preparation of cells being enumerated (e.g. Fresh NIH3T3 vs Post-Thaw NIH3T3)
concentration_type | NO | Label identifying a particular type of cell count | Character String (e.g. Total Cell Concentration) |  | The concentration type refers to the type of cell count being conducted (e.g. total cell concentration, viable cell concentration, WBC concentration, RBC concentration)
stock_solution | NO | Numerical value identifying the stock solution from which a sample has originated. | Integer (e.g. 1) | 1 | In the most simple case, all samples will originate from the same stock solution. In this case, the value in "stock_solution" will be constant.
starting_soln_conc | NO | Measured concentration of stock solution (cells/mL) | Rational Number (e.g. 235450) |  | The concentration of the stock solution should be evaluated using the same method as the one being evaluated. If more than one method is being evaluated, one of the methods should be selected, and the choice noted.
stock_extraction | NO | Numerical value identifying the order in which samples were independently prepared from the stock solution | Integer (e.g.1) |  | The experimental design specifies that samples should be generated independently from the stock solution. This numerical value indicates the order in which those samples were generated from the stock solution and can be useful for monitoring depletion effects in the stock solutions
replicate_sample | NO | Numerical value identifying the particular sample replicate | Integer (e.g. 1) |  | The experimental design specifies that replicate samples should be evaluated at each target DF. This numerical values identifies which replicate sample the data corresponds to.
analyst | NO | Numerical value identifying the analyst who obtained the observation | Integer (e.g.1) | 1 | If multiple analysts are collecting observations, this numerical value may be used to distinguish between data obtained by different analysts. If only one analyst collected the data, values in this column will be constant

## Simple Data Template for COMET
Template Purpose: This template accepts a minimal set of data in a simple format, that can be uploaded to the Cell Counting Shiny App. The Shiny App will translate the data into a format that can be analyzed be the app. The new data file can be downloaded in order to add more meta-data (e.g. measured DF, sample extraction order, time elapsed, etc.) if desired.
Instructions: Add data to the template where appropriate, leave all other fields blank. Method labels, e.g. "Method 1, Method 2...Method 6" can be replaced with more descriptive cell counting method label information. Analysis conducted directly from this template assumes: 1) the user-defined criteria for pre-evaluation of dilution integrity has been met (i.e. measured dilution fraction will not be used in the data analysis) and 2) a single stock cell solution was used to generate all samples. Analysis using this template does not include monitoring of experiment integrity (e.g. sample stability over time, pipetting error, etc.).
Method 1 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)
Method 2 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)
Method 3 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)
Method 4 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)
Method 5 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)
Method 6 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)

## COMET Input Parameters
User Interface Access
User Interface: Analyze Cell Counting Methods
Choose CSV File | Files must be *.CSV files and follow the data template format
How are missing values denoted in the file? | Default Value: "NA"
This character string should be changed to match the character string in the data template that denotes missing values
Variance is… | Defalut Value: Proportional to mean
This input corresponds to the proportional model fit. The default "proportional to mean" follows a Quasi-Poisson assumption.TThe constant mean-variance relationship follows ordinary least squares.
Order of smoothing polynomial | Default Value
This input corresponds to Calculation of the smoothed residual. The number of coefficients to be fit in the model (or order of the smoothing polynomial) should be the number of unique target dilution fractions minus 1. Generally, the order of the smoothing polynomial should not exceed the number of unique target dilution fractions, regardless of the number of samples and replicate measurements taken at each dilution fraction.
Number of Bootstrap Iterations | Input: "50 (test run)", "1000 (final run)" or "custom"
To assess the significance of the differences between PI values across cell counting measurement processes when a single evaluation of PI is performed, a non-parametric bootstrap analysis method may be applied to confidence intervals for the PI values. This value indicates the number of iterations the user wishes to run for the non-parametric bootstrap analysis. The higher the number of iterations, the better the confidence in the bootstratp analysis but the longer it takes for the program to run.
Confidence Level | Default Value: "95%"
This input allows the user to set the level of confidence to evaluate when comparing metrics via non-parameteric bootstrap analysis.
Proportionality Indicies | Select different calculations of the smoothed residual
This input selects the different types of PI that will be displayed by the program.

## COMET Output Tabs
Data Overview | Title | Indicating the Cell Type and the Concentration Type
| Raw Data | plot displays the (non-aggregated) raw data. Cell concentration is displayed as a function of the target (or measured, if present) dilution fraction.
| Mean Cell Concentration vs. Dilution Fraction | plot displays measured (or target) dilution fraction on the x-axis and cell concentration on the y-axis. The different panels indicate the different counting methods. Each point represents an average of the observations for a particular replicate sample. The black diagonal line represents the fitted proportional model. The 'Toggle Flexible Model' button toggles prediction intervals computed using the fitted flexible (polynomial) model. Prediction intervals that don't overlap with the proportional model suggests potential non-linearity.
Toggle Flexible Model - shows error bars associated with the proportional model fit
| Residual Plot for Model Fits | The top row of the plot shows the raw residuals for the proportional model. The second row shows the smoothed residuals, which indicate the difference between the flexible model's fitted values and the proportional model's fitted values. The bottom row plots the smoothed residuals (row 2), which have been scaled by the dilution fraction (x-axis).
Metrics Plots | Metrics | R2 values are plotted for each method with error bars representing confidence intervals from the bootstrap analysis.
Selected PI metrics are plotted for each method with error bars representing confidence intervals from the bootstrap analysis.
| Mean Cell Concentration vs. Dilution Fraction | Plot of average dilution fraction is given for each method, at each dilution fraction. The vertical bars represent the mean plus/minus 1 standard error of the mean (standard deviation / sqrt(# replicate samples)).
| Coefficient of Variation (CV) at each Dilution Fraction | Plot of mean %CV is computed for each method, at each dilution fraction. The vertical bars represent the mean percent CV plus minus the standard error of the mean CV (standard deviation / sqrt(# replicate samples)).
| Proportionality Constants | Plot shows the proportionality constant estimates for each counting method, along with bootstrap confidence intervals. Non-overlapping intervals suggests that the proportionality constants are different for the counting methods. Note that this difference does not suggest one method is more accurate than the other, only a difference in their proportional fits.
Metrics Tables | Mean cell concentration for each dilution fraction | Table view of mean cell concentration with upper and lower confidence intervals across replicate samples
| Mean %CV for each dilution fraction | Table view of mean % CV with upper and lower confidence intervals from bootstrap analysis
| R2 for proportional model fit | Table view of R2 with upper and lower confidence intervals from bootstrap analysis
| Proportionality Indices for model fit | Table view of proportionality index with upper and lower confidence intervals from bootstrap analysis
Experimental Design | Methods and Target Dilution Fractions | Summary of the experimental design as reconstructed from the data uploaded in the data template. This information is helpful for verifying that the data was properly entered into the GUI.
Pipetting Error
Number of Replicate Samples
Number of Replicate Observations per Sample
Stats Analysis | Statistical Modelling Details | This tab gives some of the mathematical details of the computed analysis. The proportional and flexible models are given that were fitted in the analysis based on user input, as well as the mean-variance relationship. Details of the proportional model (i.e. estimates of the proportionality constant). Equations are given for calculating the selected PIs as well as the number of bootstrap iterations.
Compare methods | Comparison Table | This tab is still under development. The comparison table provides the ratio of metrics between two methods along with the upper and lower confidence intervals. If the confidence intervals contain the value "1" (a value of "1" indicating equality) then there is no statistically significant difference between the methods ("significant" column).
| Method Precision Plots | This plot uses the fitted flexible model to estimate the range of instrument readings that could plausibly be seen at a given true dilution fraction. For example, at the x-axis value of 0.50, the vertical range on the y-axis is an estimate of the typical range expected for a replicate observation measurement for the given instrument. The range is computed by gathering all target (or measured) DF values whose prediction intervals contain the cell count corresponding to the desired dilution fraction.
| Bias comparison table | The bias comparison table provides the% bias between two methods along with the upper and lower confidence intervals. If the confidence intervals contain the value "1" indicating that the two methods could be equivalent, then there is no observed bias between the methods ("significant" column). Percent bias is calculated as 100(1−β^1β^2) , where β^1 and β^2 are the estimated proportionality constants for the two methods in the given row.
Experiment Integrity | Sample Integrity over time | Provides a visual representation of the difference from expected concentration for each method over time. Methods can be viewed individually, on their own time line, using "which method"
| Pipetting Error | Presents target dilution fraction vs measured dilution fraction with a proportional model fit.
| Viability over time comparison | Provides visual of % viability over time with linear model fits to identify any trends in changes in viability over time which can indicate a sample stability issue.
Viability Analysis | Viability Comparison | Distribution of % cell viability across samples (regardless of DF, all samples from a given method should have the same viability). Wider distributions indicate greater variability in the cell viability measurement.
| Emprical CDF and Histogram | Empirical CDF plots the absolute residual about the median for each fraction of the data for each method. A method which approaches 1 at a lower absolute residual (approaches y =1 more quickly) indicates a narrower distribution in the cell viability measurements. Similarly, the histogram represents a different visualization of the residual about the median viability for each fraction of the data for each method.
| Pairwise Kolmogorov-Smirnov Test | The KS test tests to see whether there appears to be significant differences between the distributions of the two groups being tested. Above, we are testing whether the absolute residuals about the median have similar or different spreads. A low p-value indicates evidence that the spreads of the two distributions are not equivalent.
| Equivalence of mean percent viable cells | Table for 2-sample t procedures for all counting method pairs. The confidence interval is estimated for the difference 'Method 1' - 'Method 2'. A low p-value indicates evidence that the mean percents of the two distributions are different.

## Download Results File Descr.
Download Results File | User specified parameters: | These outputs allow users to track their study design and analysis parameters used in the data analysis
File name:
var_func
Flexible model was polynomial of order
Number of bootstrap iterations conducted:
Confidence Level:
Factor Levels
stock_solution:
starting_soln_conc:
counting_method:
cell_type:
concentration_type:
random_sample_number:
stock_extraction:
target_dilution_fraction:
measured_dilution_fraction:
replicate_sample:
rep_obsv:
analyst:
Metrics
| These values will allow users to re-create plots presented in the output of the application. | counting_method | cell_type | concentration_type | Metric | Value | upper | lower
Prop.Const.x
R.squared
Scaled.Sum.of.Squares
Scaled.Absolute.Value
Absolute.Value
Smoothed.R.squared
Smoothed.Scaled.Sum.of.Squares
Smoothed.Scaled.Abs..Value
Smoothed.Absolute.Value
pooled_cv_0.1 (at each Target DF)
mean_conc_0.1 (at each Target DF)
Comparing Metrics
| These values will allow users to re-create plots presented in the output of the application | Metric | level1 (Method 1) | level2 (Method 2) | Ratio (between level 1 and 2) | lower | upper
Prop.Const.x
R.squared
Scaled.Sum.of.Squares
Scaled.Absolute.Value
Absolute.Value
Smoothed.R.squared
Smoothed.Scaled.Sum.of.Squares
Smoothed.Scaled.Abs..Value
Smoothed.Absolute.Value
pooled_cv_0.1 (at each Target DF)
mean_conc_0.1 (at each Target DF)

## Full Template Example 1 COMET
counting_method | target_dilution_fraction | measured_dilution_fraction | random_sample_number | rep_obsv | cell_conc | raw_count | percent_viable_cells | time_elapsed | cell_type | concentration_type | stock_solution | starting_soln_conc | stock_extraction | replicate_sample | analyst
Method 2 | 0.3 | 0.30034 | 1 | 1 | 278438.85 | 234 | NA | 54 | Cell Type A | Total Cell Concentration | 1 | 923700 | 6 | 3 | 2
Method 2 | 0.3 | 0.30034 | 1 | 2 | 328397.08 | 276 | NA | 54 | Cell Type A | Total Cell Concentration | 1 | 923700 | 6 | 3 | 2
Method 2 | 0.3 | 0.30034 | 1 | 3 | 323639.17 | 272 | NA | 54 | Cell Type A | Total Cell Concentration | 1 | 923700 | 6 | 3 | 2
Method 2 | 0.9 | 0.89928 | 2 | 1 | 960012.12 | 807 | NA | 55 | Cell Type A | Total Cell Concentration | 1 | 923700 | 15 | 3 | 2
Method 2 | 0.9 | 0.89928 | 2 | 2 | 960012.12 | 807 | NA | 55 | Cell Type A | Total Cell Concentration | 1 | 923700 | 15 | 3 | 2
Method 2 | 0.9 | 0.89928 | 2 | 3 | 902917.01 | 759 | NA | 55 | Cell Type A | Total Cell Concentration | 1 | 923700 | 15 | 3 | 2
Method 2 | 0.1 | 0.1013 | 3 | 1 | 101205.988 | 85 | NA | 56 | Cell Type A | Total Cell Concentration | 1 | 923700 | 2 | 2 | 2
Method 2 | 0.1 | 0.1013 | 3 | 2 | 119048.221 | 100 | NA | 56 | Cell Type A | Total Cell Concentration | 1 | 923700 | 2 | 2 | 2
Method 2 | 0.1 | 0.1013 | 3 | 3 | 97547.54 | 82 | NA | 59 | Cell Type A | Total Cell Concentration | 1 | 923700 | 2 | 2 | 2
Method 2 | 0.5 | 0.50242 | 4 | 1 | 493735.118 | 415 | NA | 59 | Cell Type A | Total Cell Concentration | 1 | 923700 | 7 | 1 | 2
Method 2 | 0.5 | 0.50242 | 4 | 2 | 508008.881 | 427 | NA | 59 | Cell Type A | Total Cell Concentration | 1 | 923700 | 7 | 1 | 2
Method 2 | 0.5 | 0.50242 | 4 | 3 | 487787.707 | 410 | NA | 60 | Cell Type A | Total Cell Concentration | 1 | 923700 | 7 | 1 | 2
Method 2 | 0.7 | 0.70034 | 5 | 1 | 691189.15 | 581 | NA | 60 | Cell Type A | Total Cell Concentration | 1 | 923700 | 10 | 1 | 2
Method 2 | 0.7 | 0.70034 | 5 | 2 | 750663.26 | 631 | NA | 61 | Cell Type A | Total Cell Concentration | 1 | 923700 | 10 | 1 | 2
Method 2 | 0.7 | 0.70034 | 5 | 3 | 744715.85 | 626 | NA | 61 | Cell Type A | Total Cell Concentration | 1 | 923700 | 10 | 1 | 2
Method 2 | 0.1 | 0.102 | 6 | 1 | 120237.71 | 101 | NA | 64 | Cell Type A | Total Cell Concentration | 1 | 923700 | 1 | 1 | 2
Method 2 | 0.1 | 0.102 | 6 | 2 | 111911.32 | 94 | NA | 64 | Cell Type A | Total Cell Concentration | 1 | 923700 | 1 | 1 | 2
Method 2 | 0.1 | 0.102 | 6 | 3 | 114290.3 | 96 | NA | 65 | Cell Type A | Total Cell Concentration | 1 | 923700 | 1 | 1 | 2
Method 2 | 0.5 | 0.4986 | 7 | 1 | 490166.68 | 412 | NA | 65 | Cell Type A | Total Cell Concentration | 1 | 923700 | 8 | 2 | 2
Method 2 | 0.5 | 0.4986 | 7 | 2 | 511577.35 | 430 | NA | 65 | Cell Type A | Total Cell Concentration | 1 | 923700 | 8 | 2 | 2
Method 2 | 0.5 | 0.4986 | 7 | 3 | 480650.83 | 404 | NA | 66 | Cell Type A | Total Cell Concentration | 1 | 923700 | 8 | 2 | 2
Method 2 | 0.9 | 0.9004 | 8 | 1 | 1006401.999 | 846 | NA | 66 | Cell Type A | Total Cell Concentration | 1 | 923700 | 13 | 1 | 2
Method 2 | 0.9 | 0.9004 | 8 | 2 | 933843.536 | 785 | NA | 66 | Cell Type A | Total Cell Concentration | 1 | 923700 | 13 | 1 | 2
Method 2 | 0.9 | 0.9004 | 8 | 3 | 929085.596 | 781 | NA | 70 | Cell Type A | Total Cell Concentration | 1 | 923700 | 13 | 1 | 2
Method 2 | 0.9 | 0.8991 | 9 | 1 | 987370.24 | 830 | NA | 70 | Cell Type A | Total Cell Concentration | 1 | 923700 | 14 | 2 | 2
Method 2 | 0.9 | 0.8991 | 9 | 2 | 917190.77 | 771 | NA | 71 | Cell Type A | Total Cell Concentration | 1 | 923700 | 14 | 2 | 2
Method 2 | 0.9 | 0.8991 | 9 | 3 | 949306.83 | 798 | NA | 71 | Cell Type A | Total Cell Concentration | 1 | 923700 | 14 | 2 | 2
Method 2 | 0.5 | 0.50053 | 10 | 1 | 479461.33 | 403 | NA | 71 | Cell Type A | Total Cell Concentration | 1 | 923700 | 9 | 3 | 2
Method 2 | 0.5 | 0.50053 | 10 | 2 | 575809.4 | 484 | NA | 72 | Cell Type A | Total Cell Concentration | 1 | 923700 | 9 | 3 | 2
Method 2 | 0.5 | 0.50053 | 10 | 3 | 515145.82 | 433 | NA | 72 | Cell Type A | Total Cell Concentration | 1 | 923700 | 9 | 3 | 2
Method 2 | 0.7 | 0.69689 | 11 | 1 | 750000 | 629 | NA | 72 | Cell Type A | Total Cell Concentration | 1 | 923700 | 12 | 3 | 2
Method 2 | 0.7 | 0.69689 | 11 | 2 | 776831.91 | 653 | NA | 105 | Cell Type A | Total Cell Concentration | 1 | 923700 | 12 | 3 | 2
Method 2 | 0.7 | 0.69689 | 11 | 3 | 736389.5 | 619 | NA | 105 | Cell Type A | Total Cell Concentration | 1 | 923700 | 12 | 3 | 2
Method 2 | 0.3 | 0.30127 | 12 | 1 | 336723.46 | 283 | NA | 105 | Cell Type A | Total Cell Concentration | 1 | 923700 | 4 | 1 | 2
Method 2 | 0.3 | 0.30127 | 12 | 2 | 350997.25 | 295 | NA | 106 | Cell Type A | Total Cell Concentration | 1 | 923700 | 4 | 1 | 2
Method 2 | 0.3 | 0.30127 | 12 | 3 | 342670.87 | 288 | NA | 106 | Cell Type A | Total Cell Concentration | 1 | 923700 | 4 | 1 | 2
Method 2 | 0.7 | 0.69631 | 13 | 1 | 794674.14 | 668 | NA | 106 | Cell Type A | Total Cell Concentration | 1 | 923700 | 11 | 2 | 2
Method 2 | 0.7 | 0.69631 | 13 | 2 | 819653.26 | 689 | NA | 107 | Cell Type A | Total Cell Concentration | 1 | 923700 | 11 | 2 | 2
Method 2 | 0.7 | 0.69631 | 13 | 3 | 753042.26 | 633 | NA | 107 | Cell Type A | Total Cell Concentration | 1 | 923700 | 11 | 2 | 2
Method 2 | 0.3 | 0.30225 | 14 | 1 | 365271.05 | 307 | NA | 107 | Cell Type A | Total Cell Concentration | 1 | 923700 | 5 | 2 | 2
Method 2 | 0.3 | 0.30225 | 14 | 2 | 289144.17 | 243 | NA | 108 | Cell Type A | Total Cell Concentration | 1 | 923700 | 5 | 2 | 2
Method 2 | 0.3 | 0.30225 | 14 | 3 | 354565.69 | 298 | NA | 108 | Cell Type A | Total Cell Concentration | 1 | 923700 | 5 | 2 | 2
Method 2 | 0.1 | 0.10469 | 15 | 1 | 103584.95 | 87 | NA | 109 | Cell Type A | Total Cell Concentration | 1 | 923700 | 3 | 3 | 2
Method 2 | 0.1 | 0.10469 | 15 | 2 | 105963.91 | 89 | NA | 109 | Cell Type A | Total Cell Concentration | 1 | 923700 | 3 | 3 | 2
Method 2 | 0.1 | 0.10469 | 15 | 3 | 101205.99 | 85 | NA | 109 | Cell Type A | Total Cell Concentration | 1 | 923700 | 3 | 3 | 2

## Full Template Example 2 COMET
stock_solution | starting_soln_conc | counting_method | cell_type | concentration_type | random_sample_number | stock_extraction | target_dilution_fraction | measured_dilution_fraction | replicate_sample | rep_obsv | analyst | time_elapsed | cell_conc | raw_count
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 1 | 2 | 4 | 271000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 2 | 2 | 7 | 346000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 3 | 2 | 10 | 328000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 1 | 2 | 13 | 753000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 2 | 2 | 15 | 829000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 3 | 2 | 18 | 808000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 1 | 2 | 21 | 123000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 2 | 2 | 24 | 152000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 3 | 2 | 26 | 113000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 1 | 2 | 29 | 490000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 2 | 2 | 32 | 505000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 3 | 2 | 46 | 543000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 1 | 2 | 49 | 650000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 2 | 2 | 51 | 657000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 3 | 2 | 54 | 661000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 1 | 2 | 57 | 118000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 2 | 2 | 60 | 121000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 3 | 2 | 62 | 121000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 1 | 2 | 65 | 613000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 2 | 2 | 68 | 600000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 3 | 2 | 71 | 531000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 1 | 2 | 73 | 833000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 2 | 2 | 76 | 815000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 3 | 2 | 79 | 760000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 1 | 2 | 82 | 797000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 2 | 2 | 84 | 908000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 3 | 2 | 87 | 774000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 1 | 2 | 90 | 449000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 2 | 2 | 93 | 559000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 3 | 2 | 96 | 477000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 1 | 2 | 98 | 611000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 2 | 2 | 101 | 611000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 3 | 2 | 104 | 672000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 1 | 2 | 112 | 379000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 2 | 2 | 114 | 393000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 3 | 2 | 117 | 347000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 1 | 2 | 120 | 622000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 2 | 2 | 123 | 626000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 3 | 2 | 125 | 565000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 1 | 2 | 128 | 295000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 2 | 2 | 131 | 349000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 3 | 2 | 134 | 362000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 1 | 2 | 136 | 117000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 2 | 2 | 139 | 121000 | NA
1 | 923700 | Method 4 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 3 | 2 | 142 | 99800 | NA
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 1 | 2 | 54 | 278438.85 | 234
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 2 | 2 | 54 | 328397.08 | 276
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 3 | 2 | 54 | 323639.17 | 272
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 1 | 2 | 55 | 960012.12 | 807
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 2 | 2 | 55 | 960012.12 | 807
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 3 | 2 | 55 | 902917.01 | 759
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 1 | 2 | 56 | 101205.988 | 85
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 2 | 2 | 56 | 119048.221 | 100
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 3 | 2 | 59 | 97547.54 | 82
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 1 | 2 | 59 | 493735.118 | 415
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 2 | 2 | 59 | 508008.881 | 427
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 3 | 2 | 60 | 487787.707 | 410
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 1 | 2 | 60 | 691189.15 | 581
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 2 | 2 | 61 | 750663.26 | 631
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 3 | 2 | 61 | 744715.85 | 626
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 1 | 2 | 64 | 120237.71 | 101
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 2 | 2 | 64 | 111911.32 | 94
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 3 | 2 | 65 | 114290.3 | 96
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 1 | 2 | 65 | 490166.68 | 412
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 2 | 2 | 65 | 511577.35 | 430
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 3 | 2 | 66 | 480650.83 | 404
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 1 | 2 | 66 | 1006401.999 | 846
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 2 | 2 | 66 | 933843.536 | 785
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 3 | 2 | 70 | 929085.596 | 781
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 1 | 2 | 70 | 987370.24 | 830
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 2 | 2 | 71 | 917190.77 | 771
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 3 | 2 | 71 | 949306.83 | 798
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 1 | 2 | 71 | 479461.33 | 403
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 2 | 2 | 72 | 575809.4 | 484
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 3 | 2 | 72 | 515145.82 | 433
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 1 | 2 | 72 | 750000 | 629
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 2 | 2 | 105 | 776831.91 | 653
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 3 | 2 | 105 | 736389.5 | 619
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 1 | 2 | 105 | 336723.46 | 283
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 2 | 2 | 106 | 350997.25 | 295
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 3 | 2 | 106 | 342670.87 | 288
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 1 | 2 | 106 | 794674.14 | 668
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 2 | 2 | 107 | 819653.26 | 689
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 3 | 2 | 107 | 753042.26 | 633
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 1 | 2 | 107 | 365271.05 | 307
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 2 | 2 | 108 | 289144.17 | 243
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 3 | 2 | 108 | 354565.69 | 298
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 1 | 2 | 109 | 103584.95 | 87
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 2 | 2 | 109 | 105963.91 | 89
1 | 923700 | Method 5 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 3 | 2 | 109 | 101205.99 | 85
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 1 | 2 | 54 | 127374.6027 | 107
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 2 | 2 | 54 | 173764.4059 | 146
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 3 | 2 | 54 | 154732.6876 | 130
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 1 | 2 | 55 | 465187.533 | 391
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 2 | 2 | 55 | 485408.7366 | 408
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 3 | 2 | 55 | 450913.7703 | 379
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 1 | 2 | 56 | 45210.3256 | 38
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 2 | 2 | 56 | 77326.3438 | 65
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 3 | 2 | 59 | 40452.396 | 34
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 1 | 2 | 59 | 240375.4128 | 202
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 2 | 2 | 59 | 245133.3387 | 206
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 3 | 2 | 60 | 236806.9572 | 199
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 1 | 2 | 60 | 329586.5787 | 277
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 2 | 2 | 61 | 386681.7189 | 325
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 3 | 2 | 61 | 352186.7229 | 296
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 1 | 2 | 64 | 68999.9698 | 58
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 2 | 2 | 64 | 41641.8774 | 35
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 3 | 2 | 65 | 58294.6291 | 49
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 1 | 2 | 65 | 248701.7793 | 209
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 2 | 2 | 65 | 236806.9572 | 199
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 3 | 2 | 66 | 228480.5907 | 192
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 1 | 2 | 66 | 510387.8809 | 429
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 2 | 2 | 66 | 502061.4697 | 422
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 3 | 2 | 70 | 484219.2663 | 407
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 1 | 2 | 70 | 488977.1772 | 411
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 2 | 2 | 71 | 490166.6773 | 412
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 3 | 2 | 71 | 462808.5924 | 389
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 1 | 2 | 71 | 1199.4822 | 1
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 2 | 2 | 72 | 230859.5462 | 194
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 3 | 2 | 72 | 282007.2902 | 237
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 1 | 2 | 72 | 261786.0867 | 220
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 2 | 2 | 105 | 397387.0708 | 334
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 3 | 2 | 105 | 379544.8376 | 319
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 1 | 2 | 105 | 178522.3318 | 150
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 2 | 2 | 106 | 179711.817 | 151
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 3 | 2 | 106 | 161869.5838 | 136
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 1 | 2 | 106 | 403334.4818 | 339
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 2 | 2 | 107 | 443776.8591 | 373
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 3 | 2 | 107 | 372407.9264 | 313
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 1 | 2 | 107 | 196364.565 | 165
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 2 | 2 | 108 | 158301.1282 | 133
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 3 | 2 | 108 | 171385.4356 | 144
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 1 | 2 | 109 | 57105.1477 | 48
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 2 | 2 | 109 | 59484.1106 | 50
1 | 923700 | Method 6 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 3 | 2 | 109 | 55915.6625 | 47
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 1 | 2 | 54 | 448534.7999 | 377
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 2 | 2 | 54 | 515145.8217 | 433
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 1 | 6 | 0.3 | 0.30034 | 3 | 3 | 2 | 54 | 512766.8215 | 431
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 1 | 2 | 55 | 1394173.129 | 1172
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 2 | 2 | 55 | 1465542.061 | 1232
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 2 | 15 | 0.9 | 0.89928 | 3 | 3 | 2 | 55 | 1369194.014 | 1151
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 1 | 2 | 56 | 139269.4248 | 117
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 2 | 2 | 56 | 193985.5946 | 163
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 3 | 2 | 0.1 | 0.1013 | 2 | 3 | 2 | 59 | 159490.6134 | 134
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 1 | 2 | 59 | 738768.4418 | 621
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 2 | 2 | 59 | 791105.671 | 665
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 4 | 7 | 0.5 | 0.50242 | 1 | 3 | 2 | 60 | 754231.7344 | 634
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 1 | 2 | 60 | 990938.7065 | 833
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 2 | 2 | 61 | 1071823.461 | 901
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 5 | 10 | 0.7 | 0.70034 | 1 | 3 | 2 | 61 | 1082528.813 | 910
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 1 | 2 | 64 | 196364.565 | 165
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 2 | 2 | 64 | 169006.4801 | 142
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 6 | 1 | 0.1 | 0.102 | 1 | 3 | 2 | 65 | 160680.0986 | 135
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 1 | 2 | 65 | 726873.6197 | 611
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 2 | 2 | 65 | 726873.6197 | 611
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 7 | 8 | 0.5 | 0.4986 | 2 | 3 | 2 | 66 | 700705.0348 | 589
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 1 | 2 | 66 | 1398931.129 | 1176
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 2 | 2 | 66 | 1307340.963 | 1099
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 8 | 13 | 0.9 | 0.9004 | 1 | 3 | 2 | 70 | 1388225.777 | 1167
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 1 | 2 | 70 | 1391794.188 | 1170
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 2 | 2 | 71 | 1278793.318 | 1075
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 9 | 14 | 0.9 | 0.8991 | 2 | 3 | 2 | 71 | 1306151.493 | 1098
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 1 | 2 | 71 | 14283.7869 | 12
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 2 | 2 | 72 | 719736.7383 | 605
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 10 | 9 | 0.5 | 0.50053 | 3 | 3 | 2 | 72 | 792295.1413 | 666
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 1 | 2 | 72 | 738768.4418 | 621
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 2 | 2 | 105 | 1034949.524 | 870
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 11 | 12 | 0.7 | 0.69689 | 3 | 3 | 2 | 105 | 1055170.758 | 887
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 1 | 2 | 105 | 449724.2702 | 378
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 2 | 2 | 106 | 460429.622 | 387
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 12 | 4 | 0.3 | 0.30127 | 1 | 3 | 2 | 106 | 483029.7662 | 406
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 1 | 2 | 106 | 1075391.872 | 904
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 2 | 2 | 107 | 1124160.631 | 945
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 13 | 11 | 0.7 | 0.69631 | 2 | 3 | 2 | 107 | 1024244.173 | 861
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 1 | 2 | 107 | 502061.4697 | 422
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 2 | 2 | 108 | 397387.0708 | 334
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 14 | 5 | 0.3 | 0.30225 | 2 | 3 | 2 | 108 | 496114.0884 | 417
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 1 | 2 | 109 | 176143.3614 | 148
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 2 | 2 | 109 | 149974.7617 | 126
1 | 923700 | Method 7 | Cell Type B | Total Cell Concentration | 15 | 3 | 0.1 | 0.10469 | 3 | 3 | 2 | 109 | 169006.4801 | 142

## Full Template Example 3 COMET
counting_method | target_dilution_fraction | measured_dilution_fraction | random_sample_number | rep_obsv | cell_conc | raw_count | percent_viable_cells | time_elapsed | cell_type | concentration_type | stock_solution | starting_soln_conc | stock_extraction | replicate_sample | analyst
Method 1 | 0.1 | 0.0976 | 15 | 1 | 255000 | NA | 83.1 | 218 | Jurkat | total_cell_conc | 1 | 2276667 | 1 | 1 | 2
Method 1 | 0.1 | 0.0976 | 15 | 2 | 268000 | NA | 76.9 | 221 | Jurkat | total_cell_conc | 1 | 2276667 | 1 | 1 | 2
Method 1 | 0.1 | 0.0976 | 15 | 3 | 265000 | NA | 81.5 | 224 | Jurkat | total_cell_conc | 1 | 2276667 | 1 | 1 | 2
Method 1 | 0.1 | 0.0989 | 17 | 1 | 271000 | NA | 81.9 | 236 | Jurkat | total_cell_conc | 1 | 2276667 | 2 | 2 | 2
Method 1 | 0.1 | 0.0989 | 17 | 2 | 303000 | NA | 81.2 | 239 | Jurkat | total_cell_conc | 1 | 2276667 | 2 | 2 | 2
Method 1 | 0.1 | 0.0989 | 17 | 3 | 290000 | NA | 76.2 | 272 | Jurkat | total_cell_conc | 1 | 2276667 | 2 | 2 | 2
Method 1 | 0.1 | 0.0955 | 12 | 1 | 269000 | NA | 82.2 | 191 | Jurkat | total_cell_conc | 1 | 2276667 | 3 | 3 | 2
Method 1 | 0.1 | 0.0955 | 12 | 2 | 278000 | NA | 76.3 | 194 | Jurkat | total_cell_conc | 1 | 2276667 | 3 | 3 | 2
Method 1 | 0.1 | 0.0955 | 12 | 3 | 274000 | NA | 78.8 | 197 | Jurkat | total_cell_conc | 1 | 2276667 | 3 | 3 | 2
Method 1 | 0.1 | 0.0995 | 13 | 1 | 325000 | NA | 76.6 | 200 | Jurkat | total_cell_conc | 1 | 2276667 | 4 | 4 | 2
Method 1 | 0.1 | 0.0995 | 13 | 2 | 326000 | NA | 81.3 | 203 | Jurkat | total_cell_conc | 1 | 2276667 | 4 | 4 | 2
Method 1 | 0.1 | 0.0995 | 13 | 3 | 283000 | NA | 79.5 | 206 | Jurkat | total_cell_conc | 1 | 2276667 | 4 | 4 | 2
Method 1 | 0.3 | 0.2969 | 2 | 1 | 824000 | NA | 76.9 | 102 | Jurkat | total_cell_conc | 1 | 2276667 | 5 | 1 | 2
Method 1 | 0.3 | 0.2969 | 2 | 2 | 848000 | NA | 72.8 | 105 | Jurkat | total_cell_conc | 1 | 2276667 | 5 | 1 | 2
Method 1 | 0.3 | 0.2969 | 2 | 3 | 789000 | NA | 75 | 108 | Jurkat | total_cell_conc | 1 | 2276667 | 5 | 1 | 2
Method 1 | 0.3 | 0.2969 | 18 | 1 | 899000 | NA | 77.8 | 245 | Jurkat | total_cell_conc | 1 | 2276667 | 6 | 2 | 2
Method 1 | 0.3 | 0.2969 | 18 | 2 | 878000 | NA | 80.3 | 248 | Jurkat | total_cell_conc | 1 | 2276667 | 6 | 2 | 2
Method 1 | 0.3 | 0.2969 | 18 | 3 | 820000 | NA | 76 | 251 | Jurkat | total_cell_conc | 1 | 2276667 | 6 | 2 | 2
Method 1 | 0.3 | 0.2976 | 1 | 1 | 904000 | NA | 72.7 | 93 | Jurkat | total_cell_conc | 1 | 2276667 | 7 | 3 | 2
Method 1 | 0.3 | 0.2976 | 1 | 2 | 812000 | NA | 75.5 | 96 | Jurkat | total_cell_conc | 1 | 2276667 | 7 | 3 | 2
Method 1 | 0.3 | 0.2976 | 1 | 3 | 786000 | NA | 73.8 | 99 | Jurkat | total_cell_conc | 1 | 2276667 | 7 | 3 | 2
Method 1 | 0.3 | 0.3015 | 19 | 1 | 811000 | NA | 78.3 | 254 | Jurkat | total_cell_conc | 1 | 2276667 | 8 | 4 | 2
Method 1 | 0.3 | 0.3015 | 19 | 2 | 847000 | NA | 75.7 | 257 | Jurkat | total_cell_conc | 1 | 2276667 | 8 | 4 | 2
Method 1 | 0.3 | 0.3015 | 19 | 3 | 826000 | NA | 79.3 | 260 | Jurkat | total_cell_conc | 1 | 2276667 | 8 | 4 | 2
Method 1 | 0.5 | 0.4996 | 7 | 1 | 1360000 | NA | 75.7 | 147 | Jurkat | total_cell_conc | 1 | 2276667 | 9 | 1 | 2
Method 1 | 0.5 | 0.4996 | 7 | 2 | 1420000 | NA | 77.5 | 150 | Jurkat | total_cell_conc | 1 | 2276667 | 9 | 1 | 2
Method 1 | 0.5 | 0.4996 | 7 | 3 | 1400000 | NA | 75.7 | 153 | Jurkat | total_cell_conc | 1 | 2276667 | 9 | 1 | 2
Method 1 | 0.5 | 0.5004 | 4 | 1 | 1400000 | NA | 76.4 | 120 | Jurkat | total_cell_conc | 1 | 2276667 | 10 | 2 | 2
Method 1 | 0.5 | 0.5004 | 4 | 2 | 1320000 | NA | 75.1 | 123 | Jurkat | total_cell_conc | 1 | 2276667 | 10 | 2 | 2
Method 1 | 0.5 | 0.5004 | 4 | 3 | 1410000 | NA | 74.5 | 126 | Jurkat | total_cell_conc | 1 | 2276667 | 10 | 2 | 2
Method 1 | 0.5 | 0.4983 | 3 | 1 | 1340000 | NA | 72.6 | 111 | Jurkat | total_cell_conc | 1 | 2276667 | 11 | 3 | 2
Method 1 | 0.5 | 0.4983 | 3 | 2 | 1320000 | NA | 73.8 | 114 | Jurkat | total_cell_conc | 1 | 2276667 | 11 | 3 | 2
Method 1 | 0.5 | 0.4983 | 3 | 3 | 1360000 | NA | 74.3 | 117 | Jurkat | total_cell_conc | 1 | 2276667 | 11 | 3 | 2
Method 1 | 0.5 | 0.5007 | 5 | 1 | 1310000 | NA | 74.7 | 129 | Jurkat | total_cell_conc | 1 | 2276667 | 12 | 4 | 2
Method 1 | 0.5 | 0.5007 | 5 | 2 | 1330000 | NA | 75.9 | 132 | Jurkat | total_cell_conc | 1 | 2276667 | 12 | 4 | 2
Method 1 | 0.5 | 0.5007 | 5 | 3 | 1340000 | NA | 74.6 | 135 | Jurkat | total_cell_conc | 1 | 2276667 | 12 | 4 | 2
Method 1 | 0.7 | 0.7003 | 10 | 1 | 1930000 | NA | 76.2 | 174 | Jurkat | total_cell_conc | 1 | 2276667 | 13 | 1 | 2
Method 1 | 0.7 | 0.7003 | 10 | 2 | 1930000 | NA | 72.5 | 177 | Jurkat | total_cell_conc | 1 | 2276667 | 13 | 1 | 2
Method 1 | 0.7 | 0.7003 | 10 | 3 | 1830000 | NA | 76 | 180 | Jurkat | total_cell_conc | 1 | 2276667 | 13 | 1 | 2
Method 1 | 0.7 | 0.7006 | 6 | 1 | 1880000 | NA | 72.9 | 138 | Jurkat | total_cell_conc | 1 | 2276667 | 14 | 2 | 2
Method 1 | 0.7 | 0.7006 | 6 | 2 | 1860000 | NA | 73.1 | 141 | Jurkat | total_cell_conc | 1 | 2276667 | 14 | 2 | 2
Method 1 | 0.7 | 0.7006 | 6 | 3 | 1870000 | NA | 72.7 | 144 | Jurkat | total_cell_conc | 1 | 2276667 | 14 | 2 | 2
Method 1 | 0.7 | 0.6995 | 11 | 1 | 1860000 | NA | 76.3 | 183 | Jurkat | total_cell_conc | 1 | 2276667 | 15 | 3 | 2
Method 1 | 0.7 | 0.6995 | 11 | 2 | 1870000 | NA | 74.9 | 186 | Jurkat | total_cell_conc | 1 | 2276667 | 15 | 3 | 2
Method 1 | 0.7 | 0.6995 | 11 | 3 | 1890000 | NA | 74.6 | 188 | Jurkat | total_cell_conc | 1 | 2276667 | 15 | 3 | 2
Method 1 | 0.7 | 0.7011 | 9 | 1 | 1860000 | NA | 75.3 | 165 | Jurkat | total_cell_conc | 1 | 2276667 | 16 | 4 | 2
Method 1 | 0.7 | 0.7011 | 9 | 2 | 2000000 | NA | 74 | 168 | Jurkat | total_cell_conc | 1 | 2276667 | 16 | 4 | 2
Method 1 | 0.7 | 0.7011 | 9 | 3 | 1900000 | NA | 76.3 | 171 | Jurkat | total_cell_conc | 1 | 2276667 | 16 | 4 | 2
Method 1 | 0.9 | 0.8995 | 14 | 1 | 2400000 | NA | 75 | 209 | Jurkat | total_cell_conc | 1 | 2276667 | 17 | 1 | 2
Method 1 | 0.9 | 0.8995 | 14 | 2 | 2360000 | NA | 75.4 | 212 | Jurkat | total_cell_conc | 1 | 2276667 | 17 | 1 | 2
Method 1 | 0.9 | 0.8995 | 14 | 3 | 2390000 | NA | 76.2 | 215 | Jurkat | total_cell_conc | 1 | 2276667 | 17 | 1 | 2
Method 1 | 0.9 | 0.9009 | 16 | 1 | 2500000 | NA | 76.8 | 227 | Jurkat | total_cell_conc | 1 | 2276667 | 18 | 2 | 2
Method 1 | 0.9 | 0.9009 | 16 | 2 | 2370000 | NA | 75.1 | 230 | Jurkat | total_cell_conc | 1 | 2276667 | 18 | 2 | 2
Method 1 | 0.9 | 0.9009 | 16 | 3 | 2540000 | NA | 74.8 | 233 | Jurkat | total_cell_conc | 1 | 2276667 | 18 | 2 | 2
Method 1 | 0.9 | 0.9011 | 8 | 1 | 2450000 | NA | 72.2 | 156 | Jurkat | total_cell_conc | 1 | 2276667 | 19 | 3 | 2
Method 1 | 0.9 | 0.9011 | 8 | 2 | 2450000 | NA | 75.1 | 159 | Jurkat | total_cell_conc | 1 | 2276667 | 19 | 3 | 2
Method 1 | 0.9 | 0.9011 | 8 | 3 | 2410000 | NA | 73.4 | 162 | Jurkat | total_cell_conc | 1 | 2276667 | 19 | 3 | 2
Method 1 | 0.9 | 0.9022 | 20 | 1 | 2460000 | NA | 76 | 263 | Jurkat | total_cell_conc | 1 | 2276667 | 20 | 4 | 2
Method 1 | 0.9 | 0.9022 | 20 | 2 | 2460000 | NA | 75.6 | 266 | Jurkat | total_cell_conc | 1 | 2276667 | 20 | 4 | 2
Method 1 | 0.9 | 0.9022 | 20 | 3 | 2440000 | NA | 72.1 | 269 | Jurkat | total_cell_conc | 1 | 2276667 | 20 | 4 | 2
Method 2 | 0.1 | 0.0976 | 15 | 1 | 156300 | NA | 32.2 | 53 | Jurkat | total_cell_conc | 1 | 2276667 | 1 | 1 | 2
Method 2 | 0.1 | 0.0976 | 15 | 2 | 156100 | NA | 43.6 | 54 | Jurkat | total_cell_conc | 1 | 2276667 | 1 | 1 | 2
Method 2 | 0.1 | 0.0976 | 15 | 3 | 183300 | NA | 42.7 | 56 | Jurkat | total_cell_conc | 1 | 2276667 | 1 | 1 | 2
Method 2 | 0.1 | 0.0989 | 17 | 1 | 196800 | NA | 38 | 60 | Jurkat | total_cell_conc | 1 | 2276667 | 2 | 2 | 2
Method 2 | 0.1 | 0.0989 | 17 | 2 | 109200 | NA | 39.1 | 62 | Jurkat | total_cell_conc | 1 | 2276667 | 2 | 2 | 2
Method 2 | 0.1 | 0.0989 | 17 | 3 | 211400 | NA | 33.8 | 63 | Jurkat | total_cell_conc | 1 | 2276667 | 2 | 2 | 2
Method 2 | 0.1 | 0.0955 | 12 | 1 | 176900 | NA | 34.4 | 45 | Jurkat | total_cell_conc | 1 | 2276667 | 3 | 3 | 2
Method 2 | 0.1 | 0.0955 | 12 | 2 | 165700 | NA | 32.4 | 46 | Jurkat | total_cell_conc | 1 | 2276667 | 3 | 3 | 2
Method 2 | 0.1 | 0.0955 | 12 | 3 | 219200 | NA | 37.5 | 46 | Jurkat | total_cell_conc | 1 | 2276667 | 3 | 3 | 2
Method 2 | 0.1 | 0.0995 | 13 | 1 | 154400 | NA | 30.1 | 48 | Jurkat | total_cell_conc | 1 | 2276667 | 4 | 4 | 2
Method 2 | 0.1 | 0.0995 | 13 | 2 | 215300 | NA | 33.1 | 48 | Jurkat | total_cell_conc | 1 | 2276667 | 4 | 4 | 2
Method 2 | 0.1 | 0.0995 | 13 | 3 | 225200 | NA | 31.6 | 49 | Jurkat | total_cell_conc | 1 | 2276667 | 4 | 4 | 2
Method 2 | 0.3 | 0.2969 | 2 | 1 | 556000 | NA | 33.8 | 13 | Jurkat | total_cell_conc | 1 | 2276667 | 5 | 1 | 2
Method 2 | 0.3 | 0.2969 | 2 | 2 | 531000 | NA | 40.7 | 14 | Jurkat | total_cell_conc | 1 | 2276667 | 5 | 1 | 2
Method 2 | 0.3 | 0.2969 | 2 | 3 | 520000 | NA | 31.3 | 15 | Jurkat | total_cell_conc | 1 | 2276667 | 5 | 1 | 2
Method 2 | 0.3 | 0.2969 | 18 | 1 | 531000 | NA | 40.1 | 64 | Jurkat | total_cell_conc | 1 | 2276667 | 6 | 2 | 2
Method 2 | 0.3 | 0.2969 | 18 | 2 | 620000 | NA | 39 | 65 | Jurkat | total_cell_conc | 1 | 2276667 | 6 | 2 | 2
Method 2 | 0.3 | 0.2969 | 18 | 3 | 422000 | NA | 30.3 | 66 | Jurkat | total_cell_conc | 1 | 2276667 | 6 | 2 | 2
Method 2 | 0.3 | 0.2976 | 1 | 1 | 579000 | NA | 36.6 | 10 | Jurkat | total_cell_conc | 1 | 2276667 | 7 | 3 | 2
Method 2 | 0.3 | 0.2976 | 1 | 2 | 495000 | NA | 34.9 | 11 | Jurkat | total_cell_conc | 1 | 2276667 | 7 | 3 | 2
Method 2 | 0.3 | 0.2976 | 1 | 3 | 560000 | NA | 33 | 12 | Jurkat | total_cell_conc | 1 | 2276667 | 7 | 3 | 2
Method 2 | 0.3 | 0.3015 | 19 | 1 | 429000 | NA | 34.7 | 67 | Jurkat | total_cell_conc | 1 | 2276667 | 8 | 4 | 2
Method 2 | 0.3 | 0.3015 | 19 | 2 | 591000 | NA | 33.7 | 68 | Jurkat | total_cell_conc | 1 | 2276667 | 8 | 4 | 2
Method 2 | 0.3 | 0.3015 | 19 | 3 | 443000 | NA | 32.1 | 69 | Jurkat | total_cell_conc | 1 | 2276667 | 8 | 4 | 2
Method 2 | 0.5 | 0.4996 | 7 | 1 | 968000 | NA | 36.4 | 29 | Jurkat | total_cell_conc | 1 | 2276667 | 9 | 1 | 2
Method 2 | 0.5 | 0.4996 | 7 | 2 | 1001000 | NA | 36.7 | 30 | Jurkat | total_cell_conc | 1 | 2276667 | 9 | 1 | 2
Method 2 | 0.5 | 0.4996 | 7 | 3 | 928000 | NA | 38.6 | 31 | Jurkat | total_cell_conc | 1 | 2276667 | 9 | 1 | 2
Method 2 | 0.5 | 0.5004 | 4 | 1 | 957000 | NA | 37.8 | 19 | Jurkat | total_cell_conc | 1 | 2276667 | 10 | 2 | 2
Method 2 | 0.5 | 0.5004 | 4 | 2 | 808000 | NA | 32.1 | 21 | Jurkat | total_cell_conc | 1 | 2276667 | 10 | 2 | 2
Method 2 | 0.5 | 0.5004 | 4 | 3 | 1024000 | NA | 38.9 | 22 | Jurkat | total_cell_conc | 1 | 2276667 | 10 | 2 | 2
Method 2 | 0.5 | 0.4983 | 3 | 1 | 895000 | NA | 33.2 | 16 | Jurkat | total_cell_conc | 1 | 2276667 | 11 | 3 | 2
Method 2 | 0.5 | 0.4983 | 3 | 2 | 920000 | NA | 34.2 | 17 | Jurkat | total_cell_conc | 1 | 2276667 | 11 | 3 | 2
Method 2 | 0.5 | 0.4983 | 3 | 3 | 1129000 | NA | 40.2 | 18 | Jurkat | total_cell_conc | 1 | 2276667 | 11 | 3 | 2
Method 2 | 0.5 | 0.5007 | 5 | 1 | 1064000 | NA | 35.9 | 23 | Jurkat | total_cell_conc | 1 | 2276667 | 12 | 4 | 2
Method 2 | 0.5 | 0.5007 | 5 | 2 | 997000 | NA | 37.6 | 24 | Jurkat | total_cell_conc | 1 | 2276667 | 12 | 4 | 2
Method 2 | 0.5 | 0.5007 | 5 | 3 | 1028000 | NA | 36.7 | 25 | Jurkat | total_cell_conc | 1 | 2276667 | 12 | 4 | 2
Method 2 | 0.7 | 0.7003 | 10 | 1 | 1210000 | NA | 33.6 | 39 | Jurkat | total_cell_conc | 1 | 2276667 | 13 | 1 | 2
Method 2 | 0.7 | 0.7003 | 10 | 2 | 1360000 | NA | 33.5 | 40 | Jurkat | total_cell_conc | 1 | 2276667 | 13 | 1 | 2
Method 2 | 0.7 | 0.7003 | 10 | 3 | 1249000 | NA | 35 | 40 | Jurkat | total_cell_conc | 1 | 2276667 | 13 | 1 | 2
Method 2 | 0.7 | 0.7006 | 6 | 1 | 1251000 | NA | 37.3 | 26 | Jurkat | total_cell_conc | 1 | 2276667 | 14 | 2 | 2
Method 2 | 0.7 | 0.7006 | 6 | 2 | 1201000 | NA | 37.1 | 27 | Jurkat | total_cell_conc | 1 | 2276667 | 14 | 2 | 2
Method 2 | 0.7 | 0.7006 | 6 | 3 | 1195000 | NA | 33.9 | 28 | Jurkat | total_cell_conc | 1 | 2276667 | 14 | 2 | 2
Method 2 | 0.7 | 0.6995 | 11 | 1 | 1432000 | NA | 35.4 | 42 | Jurkat | total_cell_conc | 1 | 2276667 | 15 | 3 | 2
Method 2 | 0.7 | 0.6995 | 11 | 2 | 1324000 | NA | 33.2 | 43 | Jurkat | total_cell_conc | 1 | 2276667 | 15 | 3 | 2
Method 2 | 0.7 | 0.6995 | 11 | 3 | 1415000 | NA | 40.4 | 44 | Jurkat | total_cell_conc | 1 | 2276667 | 15 | 3 | 2
Method 2 | 0.7 | 0.7011 | 9 | 1 | 1269000 | NA | 40.7 | 35 | Jurkat | total_cell_conc | 1 | 2276667 | 16 | 4 | 2
Method 2 | 0.7 | 0.7011 | 9 | 2 | 1244000 | NA | 36.2 | 36 | Jurkat | total_cell_conc | 1 | 2276667 | 16 | 4 | 2
Method 2 | 0.7 | 0.7011 | 9 | 3 | 1481000 | NA | 38 | 37 | Jurkat | total_cell_conc | 1 | 2276667 | 16 | 4 | 2
Method 2 | 0.9 | 0.8995 | 14 | 1 | 1567000 | NA | 38.7 | 50 | Jurkat | total_cell_conc | 1 | 2276667 | 17 | 1 | 2
Method 2 | 0.9 | 0.8995 | 14 | 2 | 1724000 | NA | 34.5 | 51 | Jurkat | total_cell_conc | 1 | 2276667 | 17 | 1 | 2
Method 2 | 0.9 | 0.8995 | 14 | 3 | 1590000 | NA | 33.3 | 52 | Jurkat | total_cell_conc | 1 | 2276667 | 17 | 1 | 2
Method 2 | 0.9 | 0.9009 | 16 | 1 | 1538000 | NA | 37.9 | 57 | Jurkat | total_cell_conc | 1 | 2276667 | 18 | 2 | 2
Method 2 | 0.9 | 0.9009 | 16 | 2 | 1622000 | NA | 35.3 | 58 | Jurkat | total_cell_conc | 1 | 2276667 | 18 | 2 | 2
Method 2 | 0.9 | 0.9009 | 16 | 3 | 1529000 | NA | 33.9 | 59 | Jurkat | total_cell_conc | 1 | 2276667 | 18 | 2 | 2
Method 2 | 0.9 | 0.9011 | 8 | 1 | 2251000 | NA | 35.6 | 32 | Jurkat | total_cell_conc | 1 | 2276667 | 19 | 3 | 2
Method 2 | 0.9 | 0.9011 | 8 | 2 | 1894000 | NA | 34.5 | 33 | Jurkat | total_cell_conc | 1 | 2276667 | 19 | 3 | 2
Method 2 | 0.9 | 0.9011 | 8 | 3 | 1578000 | NA | 36 | 34 | Jurkat | total_cell_conc | 1 | 2276667 | 19 | 3 | 2
Method 2 | 0.9 | 0.9022 | 20 | 1 | 1588000 | NA | 35.8 | 70 | Jurkat | total_cell_conc | 1 | 2276667 | 20 | 4 | 2
Method 2 | 0.9 | 0.9022 | 20 | 2 | 1509000 | NA | 35 | 71 | Jurkat | total_cell_conc | 1 | 2276667 | 20 | 4 | 2
Method 2 | 0.9 | 0.9022 | 20 | 3 | 1478000 | NA | 32.5 | 72 | Jurkat | total_cell_conc | 1 | 2276667 | 20 | 4 | 2
Method 3 | 0.1 | 0.0976 | 15 | 1 | 149795.7331 | NA | 51.5 | 103 | Jurkat | total_cell_conc | 1 | 2276667 | 1 | 1 | 2
Method 3 | 0.1 | 0.0976 | 15 | 2 | 202451.2029 | NA | 50 | 103 | Jurkat | total_cell_conc | 1 | 2276667 | 1 | 1 | 2
Method 3 | 0.1 | 0.0976 | 15 | 3 | 196550.1589 | NA | 52.9 | 104 | Jurkat | total_cell_conc | 1 | 2276667 | 1 | 1 | 2
Method 3 | 0.1 | 0.0989 | 17 | 1 | 228325.0113 | NA | 53.7 | 106 | Jurkat | total_cell_conc | 1 | 2276667 | 2 | 2 | 2
Method 3 | 0.1 | 0.0989 | 17 | 2 | 224239.6732 | NA | 53.6 | 107 | Jurkat | total_cell_conc | 1 | 2276667 | 2 | 2 | 2
Method 3 | 0.1 | 0.0989 | 17 | 3 | 232864.276 | NA | 57.7 | 107 | Jurkat | total_cell_conc | 1 | 2276667 | 2 | 2 | 2
Method 3 | 0.1 | 0.0955 | 12 | 1 | 191556.9678 | NA | 54.3 | 99 | Jurkat | total_cell_conc | 1 | 2276667 | 3 | 3 | 2
Method 3 | 0.1 | 0.0955 | 12 | 2 | 229232.8643 | NA | 52.3 | 100 | Jurkat | total_cell_conc | 1 | 2276667 | 3 | 3 | 2
Method 3 | 0.1 | 0.0955 | 12 | 3 | 186109.8502 | NA | 56.6 | 100 | Jurkat | total_cell_conc | 1 | 2276667 | 3 | 3 | 2
Method 3 | 0.1 | 0.0995 | 13 | 1 | 163413.527 | NA | 51.4 | 100 | Jurkat | total_cell_conc | 1 | 2276667 | 4 | 4 | 2
Method 3 | 0.1 | 0.0995 | 13 | 2 | 189741.2619 | NA | 52.9 | 100 | Jurkat | total_cell_conc | 1 | 2276667 | 4 | 4 | 2
Method 3 | 0.1 | 0.0995 | 13 | 3 | 173399.9092 | NA | 45.5 | 101 | Jurkat | total_cell_conc | 1 | 2276667 | 4 | 4 | 2
Method 3 | 0.3 | 0.2969 | 2 | 1 | 757149.3418 | NA | 45.2 | 82 | Jurkat | total_cell_conc | 1 | 2276667 | 5 | 1 | 2
Method 3 | 0.3 | 0.2969 | 2 | 2 | 742623.695 | NA | 43.9 | 82 | Jurkat | total_cell_conc | 1 | 2276667 | 5 | 1 | 2
Method 3 | 0.3 | 0.2969 | 2 | 3 | 760326.8271 | NA | 44.2 | 82 | Jurkat | total_cell_conc | 1 | 2276667 | 5 | 1 | 2
Method 3 | 0.3 | 0.2969 | 18 | 1 | 713572.4013 | NA | 56.4 | 107 | Jurkat | total_cell_conc | 1 | 2276667 | 6 | 2 | 2
Method 3 | 0.3 | 0.2969 | 18 | 2 | 797094.8706 | NA | 57.7 | 107 | Jurkat | total_cell_conc | 1 | 2276667 | 6 | 2 | 2
Method 3 | 0.3 | 0.2969 | 18 | 3 | 670449.3872 | NA | 58.6 | 108 | Jurkat | total_cell_conc | 1 | 2276667 | 6 | 2 | 2
Method 3 | 0.3 | 0.2976 | 1 | 1 | 649114.8434 | NA | 46.1 | 81 | Jurkat | total_cell_conc | 1 | 2276667 | 7 | 3 | 2
Method 3 | 0.3 | 0.2976 | 1 | 2 | 660463.005 | NA | 42.2 | 81 | Jurkat | total_cell_conc | 1 | 2276667 | 7 | 3 | 2
Method 3 | 0.3 | 0.2976 | 1 | 3 | 694507.4898 | NA | 44.4 | 81 | Jurkat | total_cell_conc | 1 | 2276667 | 7 | 3 | 2
Method 3 | 0.3 | 0.3015 | 19 | 1 | 422151.6114 | NA | 56 | 108 | Jurkat | total_cell_conc | 1 | 2276667 | 8 | 4 | 2
Method 3 | 0.3 | 0.3015 | 19 | 2 | 494779.8457 | NA | 60.5 | 108 | Jurkat | total_cell_conc | 1 | 2276667 | 8 | 4 | 2
Method 3 | 0.3 | 0.3015 | 19 | 3 | 586472.9914 | NA | 40.5 | 112 | Jurkat | total_cell_conc | 1 | 2276667 | 8 | 4 | 2
Method 3 | 0.5 | 0.4996 | 7 | 1 | 1074443.94 | NA | 49.6 | 89 | Jurkat | total_cell_conc | 1 | 2276667 | 9 | 1 | 2
Method 3 | 0.5 | 0.4996 | 7 | 2 | 1094416.704 | NA | 50.1 | 90 | Jurkat | total_cell_conc | 1 | 2276667 | 9 | 1 | 2
Method 3 | 0.5 | 0.4996 | 7 | 3 | 1075805.719 | NA | 47.5 | 90 | Jurkat | total_cell_conc | 1 | 2276667 | 9 | 1 | 2
Method 3 | 0.5 | 0.5004 | 4 | 1 | 733999.0921 | NA | 42.4 | 85 | Jurkat | total_cell_conc | 1 | 2276667 | 10 | 2 | 2
Method 3 | 0.5 | 0.5004 | 4 | 2 | 969586.9269 | NA | 46 | 86 | Jurkat | total_cell_conc | 1 | 2276667 | 10 | 2 | 2
Method 3 | 0.5 | 0.5004 | 4 | 3 | 920562.8688 | NA | 44.5 | 86 | Jurkat | total_cell_conc | 1 | 2276667 | 10 | 2 | 2
Method 3 | 0.5 | 0.4983 | 3 | 1 | 1158420.336 | NA | 44.2 | 83 | Jurkat | total_cell_conc | 1 | 2276667 | 11 | 3 | 2
Method 3 | 0.5 | 0.4983 | 3 | 2 | 1102587.381 | NA | 46.2 | 83 | Jurkat | total_cell_conc | 1 | 2276667 | 11 | 3 | 2
Method 3 | 0.5 | 0.4983 | 3 | 3 | 1102133.454 | NA | 46.8 | 83 | Jurkat | total_cell_conc | 1 | 2276667 | 11 | 3 | 2
Method 3 | 0.5 | 0.5007 | 5 | 1 | 884702.6782 | NA | 43.1 | 86 | Jurkat | total_cell_conc | 1 | 2276667 | 12 | 4 | 2
Method 3 | 0.5 | 0.5007 | 5 | 2 | 935542.4421 | NA | 45.1 | 86 | Jurkat | total_cell_conc | 1 | 2276667 | 12 | 4 | 2
Method 3 | 0.5 | 0.5007 | 5 | 3 | 917839.31 | NA | 45.2 | 86 | Jurkat | total_cell_conc | 1 | 2276667 | 12 | 4 | 2
Method 3 | 0.7 | 0.7003 | 10 | 1 | 1163413.527 | NA | 48 | 93 | Jurkat | total_cell_conc | 1 | 2276667 | 13 | 1 | 2
Method 3 | 0.7 | 0.7003 | 10 | 2 | 1278710.849 | NA | 46.1 | 94 | Jurkat | total_cell_conc | 1 | 2276667 | 13 | 1 | 2
Method 3 | 0.7 | 0.7003 | 10 | 3 | 1406718.112 | NA | 47.3 | 94 | Jurkat | total_cell_conc | 1 | 2276667 | 13 | 1 | 2
Method 3 | 0.7 | 0.7006 | 6 | 1 | 1046754.426 | NA | 39.1 | 87 | Jurkat | total_cell_conc | 1 | 2276667 | 14 | 2 | 2
Method 3 | 0.7 | 0.7006 | 6 | 2 | 1338629.142 | NA | 40.5 | 89 | Jurkat | total_cell_conc | 1 | 2276667 | 14 | 2 | 2
Method 3 | 0.7 | 0.7006 | 6 | 3 | 1397185.656 | NA | 42.9 | 89 | Jurkat | total_cell_conc | 1 | 2276667 | 14 | 2 | 2
Method 3 | 0.7 | 0.6995 | 11 | 1 | 1170676.35 | NA | 43 | 94 | Jurkat | total_cell_conc | 1 | 2276667 | 15 | 3 | 2
Method 3 | 0.7 | 0.6995 | 11 | 2 | 1216068.997 | NA | 41.4 | 95 | Jurkat | total_cell_conc | 1 | 2276667 | 15 | 3 | 2
Method 3 | 0.7 | 0.6995 | 11 | 3 | 1332274.172 | NA | 44.9 | 99 | Jurkat | total_cell_conc | 1 | 2276667 | 15 | 3 | 2
Method 3 | 0.7 | 0.7011 | 9 | 1 | 1070358.602 | NA | 37 | 92 | Jurkat | total_cell_conc | 1 | 2276667 | 16 | 4 | 2
Method 3 | 0.7 | 0.7011 | 9 | 2 | 1205628.688 | NA | 43.9 | 93 | Jurkat | total_cell_conc | 1 | 2276667 | 16 | 4 | 2
Method 3 | 0.7 | 0.7011 | 9 | 3 | 1280526.555 | NA | 43.2 | 93 | Jurkat | total_cell_conc | 1 | 2276667 | 16 | 4 | 2
Method 3 | 0.9 | 0.8995 | 14 | 1 | 1399909.215 | NA | 43.7 | 101 | Jurkat | total_cell_conc | 1 | 2276667 | 17 | 1 | 2
Method 3 | 0.9 | 0.8995 | 14 | 2 | 1518837.948 | NA | 47.8 | 103 | Jurkat | total_cell_conc | 1 | 2276667 | 17 | 1 | 2
Method 3 | 0.9 | 0.8995 | 14 | 3 | 1379028.597 | NA | 48.1 | 103 | Jurkat | total_cell_conc | 1 | 2276667 | 17 | 1 | 2
Method 3 | 0.9 | 0.9009 | 16 | 1 | 1472991.375 | NA | 41.6 | 104 | Jurkat | total_cell_conc | 1 | 2276667 | 18 | 2 | 2
Method 3 | 0.9 | 0.9009 | 16 | 2 | 1701770.313 | NA | 45.5 | 104 | Jurkat | total_cell_conc | 1 | 2276667 | 18 | 2 | 2
Method 3 | 0.9 | 0.9009 | 16 | 3 | 1579210.168 | NA | 41.7 | 105 | Jurkat | total_cell_conc | 1 | 2276667 | 18 | 2 | 2
Method 3 | 0.9 | 0.9011 | 8 | 1 | 1624602.814 | NA | 43 | 90 | Jurkat | total_cell_conc | 1 | 2276667 | 19 | 3 | 2
Method 3 | 0.9 | 0.9011 | 8 | 2 | 1629596.005 | NA | 45 | 91 | Jurkat | total_cell_conc | 1 | 2276667 | 19 | 3 | 2
Method 3 | 0.9 | 0.9011 | 8 | 3 | 1438946.891 | NA | 42.9 | 91 | Jurkat | total_cell_conc | 1 | 2276667 | 19 | 3 | 2
Method 3 | 0.9 | 0.9022 | 20 | 1 | 1646391.285 | NA | 37.2 | 112 | Jurkat | total_cell_conc | 1 | 2276667 | 20 | 4 | 2
Method 3 | 0.9 | 0.9022 | 20 | 2 | 2016795.279 | NA | 36.5 | 113 | Jurkat | total_cell_conc | 1 | 2276667 | 20 | 4 | 2
Method 3 | 0.9 | 0.9022 | 20 | 3 | 2024058.103 | NA | 37 | 113 | Jurkat | total_cell_conc | 1 | 2276667 | 20 | 4 | 2

## Simple Template Example 1 COMET
Template Purpose: This template accepts a minimal set of data in a simple format, that can be uploaded to the Cell Counting Shiny App. The Shiny App will translate the data into a format that can be analyzed be the app. The new data file can be downloaded in order to add more meta-data (e.g. measured DF, sample extraction order, time elapsed, etc.) if desired.
Instructions: Add data to the template where appropriate, leave all other fields blank. Method labels, e.g. "Method 1, Method 2...Method 6" can be replaced with more descriptive cell counting method label information. Analysis conducted directly from this template assumes: 1) the user-defined criteria for pre-evaluation of dilution integrity has been met (i.e. measured dilution fraction will not be used in the data analysis) and 2) a single stock cell solution was used to generate all samples. Analysis using this template does not include monitoring of experiment integrity (e.g. sample stability over time, pipetting error, etc.).
Method 1 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)
1 | 1170000 | 1360000 | 1430000 |  |  | 1190000 | 1310000 | 1330000
0.7 | 912000 | 861000 | 891000 |  |  | 909000 | 939000 | 916000
0.5 | 715000 | 669000 | 686000 |  |  | 608000 | 635000 | 681000
0.3 | 389000 | 349000 | 387000 |  |  | 415000 | 417000 | 408000
0.1 | 123000 | 133000 | 123000 |  |  | 130000 | 132000 | 134000
0.05 | 78200 | 62500 | 52900 |  |  | 53500 | 51400 | 47400
Method 2 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)
1 | 1300000 | 1310000 | 1460000 |  |  | 1350000 | 1280000 | 1420000
0.7 | 984000 | 882000 | 849000 |  |  | 989000 | 1010000 | 941000
0.5 | 767000 | 724000 | 569000 |  |  | 611000 | 680000 | 688000
0.3 | 392000 | 351000 | 323000 |  |  | 479000 | 298000 | 418000
0.1 | 140000 | 104000 | 102000 |  |  | 91400 | 109000 | 104000
0.05 | 111000 | 69700 | 70100 |  |  | 139000 | 140000 | 70100
Method 3 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)
1 | 1233533 | 1323353 | 1302626 |  |  | 1216029 | 1342699 | 1359742
0.7 | 963150.6 | 1027637 | 899124.8 |  |  | 889451.9 | 829571.6 | 915707
0.5 | 660525.1 | 617687.7 | 601105.5 |  |  | 624136.3 | 657761.4 | 678949.8
0.3 | 395209.6 | 354214.6 | 345000 |  |  | 411000 | 361123.9 | 345002.3
0.1 | 115614.9 | 125000 | 146000 |  |  | 98100 | 120221.1 | 115154.3
0.05 | 35900 | 49300 | 51589.13 |  |  | 38691.85 | 49746.66 | 45600
Method 4 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)
Method 5 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)
Method 6 | Replicate Sample 1 |  |  |  |  | Replicate Sample 2 |  |  |  |  | Replicate Sample 3 |  |  |  |  | Replicate Sample 4 |  |  |  |  | Replicate Sample 5
Target Dilution Fraction (e.g. 0.1, 0.3, 0.5, 0.7, 0.9, 1) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml) | Observation 1 (cells/ml) | Observation 2 (cells/ml) | Observation 3 (cells/ml) | Observation 4 (cells/ml) | Observation 5 (cells/ml)